Updating the firmware for your HD-T2 

1) Before you start this process you must make sure your meter has least 50% charge
2) If this is the first time you have connected your meter to your PC you will need to download drivers for your meter
3) This should be done automatically and may take some time
4) However, a manual driver installation package is available from our web site (file name CDM20814.zip at: http://download.horizonhge.com/download.asp?calltype=T2UPGRADE). To use this, do not connect the meter. Save the ZIP folder to the desktop, unzip the file and double click the filename 
5) Attach the meter via USB to your PC, and “USB Mode” will light up on the meter

6) Once this is done you will need to download the firmware update file from:
http://download.horizonhge.com/download.asp?calltype=T2UPGRADE
7) Near the bottom of the HD-T2 Software page, right-click on the Firmware link in red

8) Save the ZIP folder to the desktop, and unzip the file to the Desktop as well
9) Attach the meter via USB to your PC, and “USB Mode” will light up on the meter

10) Double click the Blue Cross icon to run the program

11) You may get a Security Warning – Unknown Publisher – in which case you should allow the program to run
12)  When the small white window opens, click the second button “Upgrade” to open a second new window titled “Upgrade Warning”. On this second window, press the “Upgrade now” button in the bottom left hand corner

13) A progress chart will scroll across the screen as the software is loaded (a few seconds) and will confirm with [Upgrade completed]

14) Unplug the meter and turn on; for the first time only of use, the loading will take about 30 seconds as the firmware unpacks

15) You have now successfully loaded your meter with the latest firmware.
Do not attempt to do an update with less than 50% battery charge.

